Decontamination and recycling of Boule’s hematology analyzers

Information for recycling authorities:

The European Directive 2012/19/EU on Waste Electric and Electronic Equipment (WEEE) aims to minimize the impact on the environment by the prevention of waste.

The Boule Swelab™ Alfa, Swelab Alfa Plus, Swelab Lumi, Medonic™ M-series, Medonic M-series M32, Medonic M51, Exigo™ H400, Exigo and Exigo EOS hematology analyzers are labeled with the WEEE symbol (as seen in the margin) and they can be subjected to waste collection and recycling of the equipment at the end of their life cycles provided that the instrument is properly decontaminated.

All plastic parts should be separated and incinerated whereas metallic parts should be autoclaved.

Information for end-users of Boule’s analyzers:

Decontamination should be done by a professional user or a trained service personnel using proper protective measures, e.g., gloves, safety goggles and appropriate laboratory clothing. The instrument should be considered as infected, and decontamination must be done before handing over to a recycler.

Decontamination is done by running the Boule Cleaning Kit/Boule EasyCleaner according to the instructions for use, including emptying of the system, followed by removal of the liquid tubing system and connected inlets, waste tube, and mixing beakers (if applicable).

Boule’s analyzers contain replaceable batteries on the electrical board. These batteries are recyclable. In the event batteries needs to be replaced, ensure that trained service personnel carry out the procedure following instructions in the relevant service manual.

If, for whatever reason, there are difficulties in following the procedure contact your local Boule representative.